The ProScan Series project takes shape from our customers’ desire to use measuring instruments capable of offering high reliability, precision and the possibility to acquire, with simplicity, any type of physical measurement  for further processing.

Introduction

The instruments have been conceived for maximum user friendliness achieved through simple and effective menus; we have taken great care in offering nothing more and nothing else than what the user can effectively use.

During the design process, we have not only taken into consideration the high quality of these instruments, but also to the simplicity in handling an heterogeneous mass of data by means of the “Galileo 2.0” program which allows the user to combine, synchronize data from different  acquisition units; for example using three “ProScan III k” acquisition units you can generate a single data file containing the readings of up to 9 temperature sensors in perfectly synchronized. Furthermore, the user can combine a ProScan III k with a ProScan II p thus obtaining temperature as well as pressure variations.

In view of the user friendly philosophy of this series of instruments, reading the manual may seem superfluous, due to the fact that the “ProScan” series is of intuitive use. However, we still recommend reading the manual thoroughly.

Furthermore, in case of doubts regarding functional characteristics of this instrument, we would be very thankful if you send us an e-mail to info@pitagora.ch.

Startup

battery installation

The “ProScan” unit uses four “AA” batteries (not included). In order to install the batteries, you must remove the back cover by unscrewing the four “cross” screws located on all four angles of the back cover.

Once the batteries are installed, the unit will power up automatically showing the message “reset” on the display and emitting a warning beep; after a few seconds, the unit is ready for operation.

Button functions

Buttons “1” and “2”, “arrow up” and “arrow down” allow you to scroll the menus.

Buttons “3” and “4”, “minus” and “plus”, allow you to modify data fields or where there is nothing to be modified, you can “enter” or “exit” a menu.

Button “5” or “Acq.” starts the data acquisition at any time.

Button “6” or “Menu” allows you to “enter” or “exit” the “Main menu” and its “submenus”.

“On/Off” button “7” allows you to turn the unit on or off at any time. The “Off” state is actually a standby condition.

Button “8” or “Enter” is used to confirm an operation or choice and also to enter menus.

Position “9” or  “Acq.” is a red light LED that flashes during data acquisition.

setting the date and time

Press the “Menu” key to enter the Main menu and use the “UP” and “DOWN” buttons to scroll the main menu until you get to “set time” for the time and “set date” for the date and press “Enter”:

1. set time: the “set time” submenu appears showing the hour on the first line, the minutes (min.) on the second line and seconds (sec.) on the third line. Using the “UP” and “DOWN” arrow buttons you can change from one line to the other, and using the “+” and “-” buttons you can increment o decrement the number or data. Holding the key down, will produce a quicker change in the number.

2. set date: the “set date” submenu similar, to the set time menu appears, showing the day, month and year on three subsequent lines. Using the “UP” and “DOWN” arrow buttons you can change from one line to the other, and using the “+” and “-” buttons you can increment o decrement the number or data. Holding the key down, will produce a quicker change in the number.

immediate use as a measurement instrument

Operation

After powering the instrument or when leaving the main menu by pressing the “Menu” or “-” button, ProScan works as a standard measurement instrument.

Probes

ProScan has three differ-ent temperature probe inputs,   (T1, T2 and T3), as printed on top of the display. The instrument will show as many temperature measurements as the number of probes inserted, or if selected, the differences between them.

- Differential options

The instrument allows you to set a differential with at least two probes to be shown on the display, that is to say that one of the two temperatures shown will be the difference in temperature between the two probes. The difference will always be referred to T3, therefore the temperature of T3 will be shown, together with the difference between T1 and T3 (T1-T3) or T2 and T3 (T2-T3). In order to use this option, you must enable the “t1-t3”  and / or “t2-t3” command by selecting it from the main menu, (the small circle on the right hand side of the display becomes full or black when enabled).

Fahrenheit and Celsius

ProScan is set from factory to acquire in degrees Celsius, however you can activate the Fahrenheit option from the main menu by going down to this option and enabling it (black circle), this will show Fahrenheit degrees on the display.

Data Acquisition Operation

In order to start the acquisition of temperature data, the user should perform the desired settings.

- Acquisition Rate setting

The “acquisition rate” or data acquisition frequency, can be set from the main menu by selecting “acq. rate” and “Enter”. The minimum acquisition rate is 1 second (the instrument will record the temperatures every second) while the maximum is 1 hour 59 minutes and 59 seconds (the instrument will record the temperatures every two hours). The data can be changed using the usual “+” or “-” buttons. Once set, “Enter” takes you back to the main menu.

- Starting the Acquisition .

Pressing the “Acq.” button will start data acquisition at any time. Once the “Acq.” button has been pressed,  the display will show information such as the filename used for this acquisition data set, (press arrow down for each screen), the acquisition rate, currently chosen, and the amount of memory left (in days or hh/mm/ss) according to the current number probes and acquisition rate selected, and finally the instrument will request the confirmation whether to start  the acquisition or not, pressing “Enter” on “yes” starts the acquisition (the red light or LED starts blinking). If the acquisition rate is not the desired one or there is not enough storage memory, select “no” and you can go back to change the settings.

- Information available on the current acquisition

In order to view the current acquisition settings while the instrument is recording, press the “Acq.” Button again. The following information will be displayed, current file name being stored, the acquisition rate, the amount of memory left, and you will be asked to confirm whether you want to stop the acquisition or not. Pressing “Enter” will abandon the menu and continue the acquisition.

- Transferring data to your PC

In order to transfer the data acquired to your PC, or if you wish to perform remote operations from it, the instrument must be connected to the PC using the serial RS-232 communications cable supplied with the Pro Scan. When the unit is connected, a short beep can be heard. The ProScan must be turned on and in the temperature measurement mode (with or without probes). If the instrument is in one of the menus, the communication session will not be started.

- Canceling the memory

In order to cancel the memory contents completely, go down to the “clear mem.” Option on the main menu and press “Enter”. The unit will ask you to confirm, select “yes” and press “Enter” again to confirm. The unit will beep while the whole memory is erased.

WARNING! Once the memory has been cancelled, all data will be definitively lost, there is no possibility to recover it, so please verify that the data has already been transferred or downloaded to the PC.

Display messages

- No probes

No probes have been inserted in the sockets (this message usually appears when the unit is powered on and no probes are connected).

- Insert probes!

Please insert a probe before you start your acquisition (this message usually appears when the acquisition mode is activated).

- Low battery!

The batteries are almost discharged and it is highly recommended to replace them. Data accuracy is not guaranteed while the instrument is in this mode of operation.

- Memory full!

The storage memory is full, therefore you cannot start a new acquisition. Please, first download the data previously stored and then cancel the memory.

- Acquisition halted!

This message appears when the memory becomes full during the acquisition.

- Stop acq. before changing!

This message appears when the unit is in acquisition mode and a probe is removed or changed from one socket to the other. The unit will enter a waiting mode in which a reset to the original status or a confirmation to stop the acquisition is given.

- Reset !

Appears when the unit is turned on for the first time or if there are remote communication problems between the unit and the PC.

- Please Wait !

Appears every time the ProScan is processing data.

- File name:

File name automatically generated by ProScan combining the instrument’s serial number and an additional  sequential number..

- Acq. Rate:

Time passed between one acquisition and the next.

- Time left:

Maximum time available for the acquisition calculated according to the remaining memory space, to the number of probes used, and to the acquisition rate.

- Not valid date !!!

An invalid date has been entered.

- Not valid acq. Rate!

The acquisition rate selected is not acceptable.

- ProScan T3k serial no.

Instrument’s serial number made up of letters and numbers.

- Deleting memory!

This message appears, together with a modulated beeping sound, while the memory is being cancelled.

- Transfer all data!

This message appears, together with a modulated beeping sound, while the data acquired is being transferred to the PC.

Specifications

Electrical

Measurement Range:

K-type thermocouple: -200°C to +1370°C (-328°F to +2498°F)

* This specification applies only to the data logger and may differ from the thermocouple range.

Resolution: 0.1°C or 0.2 °F

Accuracy: Specified for operating temperatures over the range of 18°C to 28°C (64°F to 82°F).

K-type thermocouple: ±0.1% of reading  ±0.3°C  ±0.1% of reading ±0.5°F

Input Protection: 60V dc or 24V rms ac maximum input voltage on any combination of input pins.

Reading and Logging Rate: from 1 second to 2 hours (setting resolution of 1 second).

Battery: 4x standard 1.5V AA alkaline batteries.

Input Connector: Accepts standard miniature thermocouple connectors (flat blades spaced 7.9 mm. or 0.312 in. center to center IEC 584-3.).

Environmental

Ambient Operating Range:-25°C to 70°C (-13°F to 158°F)

Storage Temperature: -40°C to 80°C (-40°F to 176°F)

Humidity: 0% to 90% (0°C to 35°C) (32°F to 95°F)

0% to 70% (35°C to 50°C) (95°F to 122°F)

Physical

External Dimensions: 150 mm. x 84 mm. x 45 mm. (5.9 in. x 3.3 in. x 1.8 in.)

Ingress protection: IP40/DIN  VDE  0470

Weight: 350 grams.(12.5 oz).

Disclaimer

During normal operation and proper use of the ProScan instruments, it is quite improbable for the data acquired to be lost. However, Pitagora S.A. declines any responsibility for the loss or corruption of data due to improper use, malfunctioning, hardware defects or for any other cause.

Pitagora S.A. reserves the right to modify the contents of this manual at any time and without prior notice.